UDP4303S DC Power Supply Firmware Upgrade Guide

- Prepare the Upgrade File

Select the appropriate upgrade file according to your instrument model and copy it to the root directory of a USB flash drive formatted in FAT32.

- Check the Current Version

Press the MENU button, navigate to the Setup interface, and select About to view the current firmware version. Ensure that the new version is higher than the existing one.

- Insert USB Flash Drive

Insert the prepared USB drive into the device’s USB port. A notification will appear once the USB drive is recognized.

- Select Upgrade File

Go to Upgrade > File, browse the D drive, locate the upgrade file, and click Select to confirm.

- Verify Version Information

The system will display both the current and target firmware versions. After verifying the information, click Start.

- Start Upgrade Process

In the pop-up confirmation window, select Yes to begin the upgrade process.

- Wait for Completion

The upgrade process will go through several stages such as Checking, Programming, and Updating.

⚠️ Do not power off the device during the upgrade.

- Finish the Upgrade

Once the message “Successfully! Please reboot the device." appears, the upgrade is complete.

Restart the instrument as instructed.

- Verify Upgrade

After restarting, go to Setup > About again to confirm the firmware version has been updated successfully.